## Slide 1: View interpolation using convolutional neural network : Comparison study
2017.02.23
Hoyeon Lee

## Slide 2: Contents
Introduction
Methods
Results
Conclusions
Future works

## Slide 3: Introduction
Sparse-view CT is very promising technique for;
  Reducing radiation exposure to patients
  Less contamination from motion
  Reducing operation time

## Slide 4: Introduction
Analytic reconstruction of sparse-view CT introduces streak artifacts in result
Fig. 1. (a) Reconstructed slice using full-view(680 views), and (b) using sparse-view(170 views)
(a)
(b)

## Slide 5: Introduction
It can be reconstructed with iterative algorithm based on compressed sensing theory
However, those algorithms are computationally expensive
Algorithm dependent artifacts occur in reconstructed images
Stopping the reconstruction is not robust

## Slide 6: Introduction
Alternatively, one can make fully-sampled sinogram from sparsely-sampled sinogram using interpolation
Through interpolation, analytic reconstruction algorithm work for reconstruction

## Slide 7: Introduction
Convolutional neural network is one of the supervised deep learning method
CNN is usually composed of activation function and three layers;
  Convolution layer
  Pooling layer
  Fully-connected layer
Notes: Deep learning application

## Slide 8: Introduction
CNN is applied various fields in image processing
  Segmentation
  Localization
  Image captioning
  Visual questioning

## Slide 9: Introduction
Slides from: Spring 2016, EE837, Deep learning for computer vision, Prof. Junmo Kim
Fig. 2. Deep learning application to segmentation
Fig. 3. Deep learning application to localization

## Slide 10: Introduction
Mateusz Malinowski, et al., “Ask your neurons: A neural-based approach to answering questions about image”, ICCV (2015)
Fig. 4. Deep learning application to image based questioning

## Slide 11
Slides from: Spring 2016, EE837, Deep learning for computer vision, Prof. Junmo Kim
Fig. 5. Deep learning application to image captioning

## Slide 12: Purpose
Purpose of this study is to interpolate sparsely-sampled sinogram to fully-sampled sinogram using CNN

## Slide 13: Methods
For simulating CT projection, we obtained projection using XCAT phantom
W. Paul Segars, et al.,”Mcat to XCAT: The Evolution of 4-D Computerized phantoms for imaging research”, Proc. IEEE, Institute of Electrical and Electronics Engineers, Vol. 97, No. 12 (2009)
Fig. 6. XCAT phantoms
Notes: Projection data / sinogram 680 views

## Slide 14: Methods
For sparse-view CT, projection data were quarterly-sampled the original sinogram
Fig. 7. (a) Full-sampled sinogram, and (b) down-sampled sinogram
Notes: Projection data / sinogram 170 views

## Slide 15: Methods
The quarterly-sampled sinogram is up-sampled again
Fig. 8. (a) Full-sampled sinogram, (b) down-sampled sinogram, and (c) up-sampled sinogram from (b)
Notes: Original/down-sample/up-sample image

## Slide 16: Methods
Our network is designed as follows
Fig. 9. Structure of CNN for training

## Slide 17: Methods
The sinogram are divided into patches for training
Fig. 10. (a) Up-sampled sinogram, and (b) extracted patches from (a)

## Slide 18: Methods
In previous study, we only used sinogram obtained from abdomen region
Fig. 11. Reconstructed image in abdomen region
Fig. 12. Volume used for training in the previous study

## Slide 19: Methods
In this study, we compared results of trained network using different dataset
  Abdomen region
  Pelvic region
  Chest region
  Whole body
Fig. 13. Volume used for current study
Notes: (7*7)

## Slide 20: Methods
The networks have same structure, with different learning rate
Learning rates did not change while training
Momentum was 0.1 and weight decay was 0.0001 for all trainings
Dataset | Learning rate
Chest region | 0.0001
Pelvic region | 0.0001
Abdomen region | 0.0005
Whole body | 0.0001
Table. Ⅰ. Learning rates for each training dataset

## Slide 21: Methods
After training was done, we tested model with following network
Fig. 14. Structure of network for testing

## Slide 22: Results
Four different trained network were tested with three sinogram from different region of body
The results from CNN are compared to interpolation methods(linear, directional)
Interpolated sinogram were reconstructed using filtered back-projection(FBP) algorithm

## Slide 23: Results
Training errors
Fig. 15. Training error using whole body sinogram
Fig. 16. Training error using chest sinogram

## Slide 24: Results
Training errors
Fig. 17. Training error using pelvic sinogram
Fig. 18. Training error using abdomen sinogram

## Slide 25: Results
Fig. 19. (a) Ground truth sinogram, (b) linear interpolated sinogram, (c) directional interpolated sinogram, (d) deep interpolated sinogram using chest dataset, and (e) deep interpolated sinogram using whole body dataset
Notes: 0002

## Slide 26: Results
Fig. 20. Expansion of red-boxed regions in (a) Ground truth sinogram, (b) linear interpolated sinogram, (c) directional interpolated sinogram, (d) deep interpolated sinogram using chest dataset, and (e) deep interpolated sinogram using whole body dataset

## Slide 27: Results
Fig. 21. Reconstructed image of (a) Ground truth sinogram, (b) linear interpolated sinogram, (c) directional interpolated sinogram, (d) deep interpolated sinogram using chest region dataset, and (e) deep interpolated sinogram using whole body dataset

## Slide 28: Results
Fig. 22. Line profiles of red-line in Fig. 21

## Slide 29: Results
 | RMSE | PSNR
Whole body dataset | 0.008940 | 40.972776
Chest region dataset | 0.010569 | 39.519523
Linear | 0.019772 | 34.079055
Directional | 0.036504 | 28.753100
Table Ⅱ. Quantitative comparison results of pelvic region sinogram
 | RMSE | PSNR
Whole body dataset | 0.000176 | 75.075560
Chest region dataset | 0.000200 | 73.969920
Linear | 0.000506 | 65.918773
Directional | 0.000896 | 60.958678
Table Ⅲ. Quantitative comparison results of pelvic region reconstructed images

## Slide 30: Results
Fig. 23. (a) Ground truth sinogram, (b) linear interpolated sinogram, (c) directional interpolated sinogram, (d) deep interpolated sinogram using abdomen dataset, and (e) deep interpolated sinogram using whole body dataset

## Slide 31: Results
Fig. 24. Expansion of red-boxed regions in (a) Ground truth sinogram, (b) linear interpolated sinogram, (c) directional interpolated sinogram, (d) deep interpolated sinogram using abdomen dataset, and (e) deep interpolated sinogram using whole body dataset

## Slide 32: Results
Fig. 25. Reconstructed image of (a) Ground truth sinogram, (b) linear interpolated sinogram, (c) directional interpolated sinogram, (d) deep interpolated sinogram using abdomen region dataset, and (e) deep interpolated sinogram using whole body dataset

## Slide 33: Results
Fig. 26. Line profiles of red-line in Fig. 27

## Slide 34: Results
 | RMSE | PSNR
Whole body dataset | 0.007089 | 42.988355
Abdomen region dataset | 0.007155 | 42.908087
Linear | 0.013066 | 37.677442
Directional | 0.028871 | 30.790824
Table Ⅳ . Quantitative comparison results of abdomen region sinogram
 | RMSE | PSNR
Whole body dataset | 0.000152 | 76.357711
Abdomen region dataset | 0.000156 | 76.126944
Linear | 0.000316 | 70.004679
Directional | 0.000675 | 63.407814
Table Ⅴ. Quantitative comparison results of abdomen region reconstructed images

## Slide 35: Results
Fig. 28. (a) Ground truth sinogram, (b) linear interpolated sinogram, (c) directional interpolated sinogram, (d) deep interpolated sinogram using pelvic dataset, and (e) deep interpolated sinogram using whole body dataset

## Slide 36
Fig. 29. Expansion of red-boxed regions in (a) Ground truth sinogram, (b) linear interpolated sinogram, (c) directional interpolated sinogram, (d) deep interpolated sinogram using pelvic dataset, and (e) deep interpolated sinogram using whole body dataset

## Slide 37: Results
Fig. 30. Reconstructed image of (a) Ground truth sinogram, (b) linear interpolated sinogram, (c) directional interpolated sinogram, (d) deep interpolated sinogram using pelvic region dataset, and (e) deep interpolated sinogram using whole body dataset

## Slide 38
Fig. 31. Line profiles of red-line in Fig. 32

## Slide 39: Results
 | RMSE | PSNR
Whole body dataset | 0.005156 | 45.753042
Pelvis region dataset | 0.005862 | 44.639281
Linear | 0.013027 | 37.703008
Directional | 0.045775 | 26.787440
Table Ⅵ. Quantitative comparison results of chest region sinogram
 | RMSE | PSNR
Whole body dataset | 0.000085 | 81.455950
Pelvis region dataset | 0.000089 | 81.022142
Linear | 0.000316 | 69.999709
Directional | 0.000869 | 61.221183
Table Ⅶ. Quantitative comparison results of chest region reconstructed images

## Slide 40
One big mistake!

## Slide 41: Conclusions & Discussions
CNN works well for interpolating sparsely-sampled sinogram
It outperforms than the other interpolation method
Review of directional interpolation required

## Slide 42: Future works
Train the network with new dataset
Use 3D convolution to perform interpolation in cone-beam CT
Compare the method to iterative reconstruction results
Add noise to the projection data to train the network robust to noise

## Slide 43: Future works
Apply deconvolution to train the network without interpolating the sinogram
Olaf Ronneberger, et al., “U-Net: Convolutional networks for biomedical image segmentation”, MICCAI, Vol. 9351, 234-241(2015)

## Slide 44
Thank you for listening

## Slide 45: Methods
Convolutional layer
Pooling layer
Fully-connected layer
Activation function

## Slide 46: Methods
Convolution layer

## Slide 47: Methods
Pooling layer

## Slide 48: Methods
Fully-connected layer

## Slide 49: Methods
Activation functions
  Introduce non-linearity in networks
By Qef (Created from scratch with gnuplot) [Public domain], via Wikimedia Commons
tanh
sigmoid
ReLU